## Slide 1: Quality Assurance and Examples from the Transportable Array (TA)
Gillian Sharer
North American Seismic Network Training Workshop, Quality Assurance
August 1-2, 2018

## Slide 2
What is quality assurance?
What does “quality assurance” mean
to you as a network operator?
Notes: Methods for ensuring that a product is “fit for purpose” -> basic research, hazard assessment, education

## Slide 3
Reasons for performing quality assurance
Find station problems in a timely manner so that they can be fixed and/or documented
Provide correct response information to maximize data utility
Mitigate environmental conditions and inform station siting, installation, and maintenance
Assess the effect of station upgrades
Document the “goodness” of the network
Determine data suitability for research objectives

## Slide 4
Categories of quality assurance
Sensor and digitizer health
Metadata accuracy
Timing quality and GPS
Availability and telemetry
Environmental conditions and noise
Compliance with data format standards

## Slide 5
Station quality examples from the USArray Transportable Array
Station examples

## Slide 6
Corrosion-related noise
BHZ
BHE
BHN
2018/04/01
2018/04/02
TA.SUSD
scale +25000/-25000 counts
A service visit on 2018/05/03 found "Damp concrete, damp bedding sand, white corrosion in plug"
The noise resolved after cleaning the sensor connectors and mating plug and drying out the vault
Notes: Water and humidity is the enemy of electronic equipment
Noise tends to be intermittent
Metrics : dead_channel_lin, num_spikes, dc_offsets

## Slide 7
Loss of differential output
BHZ
BHE
BHN
2017/12/03
2017/12/09
TA.S22A
scale +5000/-5000 counts
BHN experienced a sudden drop in DC offset of -2500 counts and began recording amplitudes ~50% lower than normal.
This can be the result of damage to the sensor, digitizer, or cables.
The sensor was replaced on 2018/06/14, fixing the problem.
Notes: Sensors often use differential signaling to minimize electromagnetic noise.
Wikipedia: "Differential signaling is a method for electrically transmitting information using two complementary signals. The technique sends the same electrical signal as a differential pair of signals, each in its own conductor... The receiving circuit responds to the electrical difference between the two signals... external electromagnetic interference tends to affect both conductors identically. Since the receiving circuit only detects the difference between the wires, the technique resists electromagnetic noise”
Metrics: dc_offset

## Slide 8
Guralp CMG-3T sensor noise - SNOFLU
LHZ
LHE
LHN
2014/02/15
2014/02/20
TA.Z58A
scale +25000/-25000 counts, bandpass filtered 30-300 seconds
Sudden
Noise
Onset
Fixed (by mass)
Lock and
Unlock
Notes: Long period noise > 10 seconds, best seen filtered but large amplitude noise can be seen unfiltered

## Slide 9
TA.L19K..BNN and BNZ have sensor noise that has been traced to the use of zinc plating material which can deform in high temperature and humidity environments (e.g., Anderson et al., 2015)
Kinematics has since switched to using nickel plating for new instruments which eliminates this issue
Kinematics EpiSensor "zinc whiskers"
BNE
BNN
BNZ
PDFs
Waveforms